LSR Quick Notes

Initial Set-up:

- Turn on the LSRII cytometer allowing ample laser warm-up time.

- Fill sheath fluid tin and empty the waste can if needed.

- Clear any air bubbles in the sheath fluid line.

- Log in to Windows using your standard  hospital login and open the DIVA software.

Special Note:  If you are using either the Behemoth or the Octopus raise the cytometer lid GENTLY.

Quality Control:

- Log on as “Quality Control”.  Password=qc.

- Scroll to last QC dated experiment and duplicate without data.

- If you get a warning window that says “CST Mismatch” check the “Don’t show this message again for the current login session” and then choose “Keep BD FACSDiva Settings.”

- Rename experiment with current date and delete the numerical suffix.

- Add 1 drop of Spherotech Rainbow beads (vortexed thoroughly) to approx. 400uL sheath fluid.

- Run beads on “Low” observing histograms, looking for stable, sharp, symmetrical peaks>Record.

Experiment set-up:

Set Up an Experiment from Scratch

- Log in under lab PI name and find personal folder.

- Create and name a new experiment.

- Add a needle icon and rename it “Voltages”.

- Rename the sample tube under “Voltages” to “Unstained Cells”.

- Go to the “Instrument” window and click on “Parameters”.

- Delete detectors that are not needed.

- Create a FSC/SSC dot plot on the standard worksheet of the graphics screen.

- Run “Unstained Cells” and adjust FSC/SSC voltages in order to see cells on the dot plot.

- Gate on your cells on the dot plot.

- Create a histogram for each color and designate its data source as the P1 gate from the dot plot.

- Check the “Mark Grid” checkbox in the Inspector for each histogram.

- Run unstained sample and adjust voltages for colors to set background > Record.

Set Up an Experiment via “Duplicate Without Data”

- Log in under lab PI name and find experiment that you wish to duplicate.

- Open and highlight the experiment name>right click> “Duplicate Without Data”.

- In the duplicated experiment “Delete” and “Unlink” the previous compensation set-up as well as all experiment sample tubes but the first.

- Activate the “Unstained Cells” tube under the “Voltages” needle icon and open the voltage graphs on the “Standard Worksheet”.

- Run unstained sample and adjust voltages for scatter and colors as needed > Record.

Compensation: You must ALWAYS create a new compensation panel if using multiple fluorochromes.

- Go to “Experiment” at the main menu bar >Compensation Set-up>Create Compensation Controls.

- Activate the “Unstained Control” sample in the compensation panel.

- Place unstained cells on the SIP running on high and “Record”.

- Repeat for each of the single stain compensation samples listed on the compensation panel.

- Adjust the P1 gate on the scatter graph around the population of interest.  Activate the P1 gate, right click and “Apply to all Compensation Controls”.

- Set interval gates for each single color sample in order to define positive staining for that color.

- After setting all gates go to “Experiment” >Calculate Compensation>Link + Save.

Running Samples:

Run Samples from an Experiment Set Up from Scratch

- Create a new needle icon (experiment group) for your samples and name it.

- Switch to and create a “Global” worksheet on the graphics screen.

- Create a dot plot for FSC/SSC with a polygon gate for the population of interest.

- Add dot plots and/or histograms specific to your experiment.

- Set “events to record” at the acquisition control window to the number desired.

- Run first sample and record.

- After 1st sample, click “Next” on the acquisition control window or “New Tube” at the top of the browser to create the next sample tube.

- Rename second sample, acquire and record and so on.

Run Samples from an Experiment Set Up Via “Duplicate Without Data”

- Click on the needle icon that you named previously for your samples and rename it if necessary.

- On the graphics screen switch to the “Global” worksheet revealing the graphs that were made in the previous experiment.

- Set “events to record” at the acquisition control window to the number desired.

- Run first sample and record.

- After 1st sample, click “Next” on the acquisition control window or “New Tube” at the top of the browser to create the next sample tube

- Rename second sample, acquire and record and so on.

Print Your Data   (if needed)

- Click on the PDF icon at the top of the graphics window.

- Save the PDF to your network drive.

- Print from your local printer.

Export Your Data:

- Highlight the experiment name that is open.

- File>Export>FCS files.

- Export as FCS 3.0 files to your folder on your U drive or other network drive that is yours.

Shut down:

- While exporting data clean the machine by putting 10% bleach on the SIP with the arm in the open position for 1 minute followed by 5 minutes with the SIP arm closed running in High.

- Repeat with DI water.

- Close experiment and log out of DIVA.

- If someone is to follow, leave the computer and the cytometer on, otherwise turn the cytometer off  but leave the computer on.

Special Notes:

If you are using either the Behemoth or the Octopus close the lids VERY GENTLY if shutting down.

ABSOLUTELY NO INTERNET SURFING OR E-MAIL USAGE ON THE CYTOMETER COMPUTERS.